Sample Requirements

- ICP-OES

To ensure accurate results and prevent damage to the instrument, samples for Inductively Coupled Plasma Optical Emission Spectroscopy (ICP-OES) must meet specific physical and chemical criteria.

- Form: Samples must be in a liquid state (aqueous). Solid samples must undergo acid digestion.

- Filtration: Solutions must be free of suspended solids to prevent clogging the nebulizer. Use a 0.45 µm (standard) or 0.2 µm (for ultra-trace) membrane filter.

- Viscous Sample: High-viscosity samples (like oils or concentrated acids) must be diluted to match the viscosity of the calibration standards to ensure consistent uptake rates.

- Acid Matrix:

Nitric Acid (HNO3) is the preferred matrix. It keeps most elements in solution and is "plasma-friendly”.

Hydrochloric Acid (HCl): Required for stabilizing elements like Ag, Au, Pd, Pt, Sb, and Sn.

Aqua regia and Reverse aqua regia can be used in some special cases.

Hydrofluoric Acid (HF): Only used if the sample contains silica (SiO2) or refractory elements.

For Dilution: It is required to use only Type I Ultrapure Water.

Note: For details about Acid Digestion consider Standard guideline USEPA 3051a & USEPA 3050 b.

- Total Dissolved Solids (TDS): TDS should be < 1% (10,000 ppm).

- Minimum Volume: Typically required 10 mL per sample solution.

- Minimum 100mg of Solid/powder sample required.

- Methods:

IntelliQuant Screening: It is a semi-quantitative method that performs a rapid, full-spectrum scan (typically 167–785 nm) of a sample.

It identifies which elements are present in your sample and their approximate concentrations without needing a dedicated standard calibration curve for every element.

Quantitative Estimation: It is the most accurate way to measure a specific element at a specific wavelength.

Standard solution (concentration known) required for every element, to form Linear Calibration Curve and target element’s concentration must be within the Linear Dynamic Range.

- FTIR

To get reliable results from Fourier Transform Infrared (FTIR) spectroscopy, your samples need to be prepared specifically to allow infrared light to pass through or reflect off them correctly.

- Solid/Powders: Only a tiny amount (2-10mg) is needed, just enough to cover the active area of the crystal (a circle 2 mm in diameter). It should be Moisture free.

- Liquids: A single drop is sufficient. The liquid must be viscous enough not to evaporate instantly or runny enough to wet the crystal surface.

- Films/Polymers: Ensure the film is not too thick to be clamped, but thick enough to be handled.

Note: Lint-free tissue and a high-purity solvent (Isopropyl alcohol) is required to clean the crystal between every run.

# Flash Point Tester

Minimum 50mL sample is required.

- Automatic Kinematic Viscometer

- Minimum Volume: Approximately 10 mL for a single measurement.

- The sample must be Uniform and Bubble free.

- Sample must be filtered through a 75 micron screen.

- A solvent is required that can completely dissolve the sample (e.g., n-Hexane, Toluene, Petroleum Ether) for the automated cleaning cycle.

- Brookfield Viscometer

- Minimum Volume: Approximately 15 mL for a single measurement.

- The sample must be Uniform and Bubble free.

- Lint-free tissue and solvent that can completely dissolve the sample is required for cleaning purposes.

- Millipore Test System

Not Required